Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

February 12, 2026

liu Arnelo Qiuping
Karolinska Huddinge
SE-Huddinge
USA

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| LF1923 | JAW OPEN LF1923 LIGASURE MARYLAND 23CM | LN: 51410230X |

708681305

MPXR Report # 1360638

Event date: 2025-11-03

Dear liu Arnelo Qiuping,

Thank you for informing Medtronic of your experience with the above referenced product
      . The information we received stated during a pancreatic resection (whipple without vessel) procedure involving the LigaSure Vessel Sealing device, the device was physically broken or damaged when a piece of metal came loose from the handle. The metal piece was found near the surgical area of the patient. The device was plugged into an FT10 generator at the time of the event. A new handpiece was picked up to complete the procedure. Photos were submitted showing metal piece detached from the handpiece. No x-ray or any medical intervention was needed. There was no patient injury. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: LF1923 Lot No: 51410230X One JAW OPEN LF1923 LIGASURE MARYLAND 23CM was received for evaluation. Examination of the sample is described below. Visual inspection: • The visual inspection found the metal clicker of the device was broken off. The broken piece was returned. Evaluation: • The investigation found the metal clicker on the device was broken and detached from the device. • The clicker tab failure is due to extensive use of the device. A large number of cycles tend to fatigue the clicker. • The device's usage data was reviewed and it was identified the device had 800+ initiated seals. The product sample was not returned to the Visual inspection: A visual inspection of the returned photo(s) noted: • The picture shows a black piece of metal laying on the body of the device. Evaluation: • Without receiving the device, a detail investigation could not be performed for the reported complaint. • The customer is advised to return the device, so a complete evaluation can be performed. • If additional information is obtained, or the product is returned, this investigation will be re-opened. The reported condition was confirmed. It was determined that the most likely cause was traced to a component failure. Medtronic's assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic's complaint monitoring and tracking system. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical