Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

February 4, 2026

Izhar Haque
120 Angas Street
Adelaide SA 5000

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| LF1923 | JAW OPEN LF1923 LIGASURE MARYLAND 23CM | LN: UNKNOWN |
| VLFT10GEN | VLFT10GEN FT SERIES ENERGY PLATFORMX1 | SN: UNKNOWN |

707556166

MPXR Report # 1269907

Event date: 2025-02-05

Dear Izhar Haque,

Thank you for informing Medtronic of your experience with the above referenced products. The information we received stated during a procedure, the handpiece had message in the unit saying not completing seal cycle. The procedure was completed by replacing the device. There was no patient injury. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: LF1923 Lot No: UNKNOWN JAW OPEN LF1923 LIGASURE MARYLAND 23CM was received for evaluation. Examination of the sample is described below. Visual inspection: - The visual inspection found seal plate damage, consistent with arcing. Evaluation: - The damage to the seal plate(s) is consistent with the user inadvertently closing jaws on a hard/metallic object and activating the device. - Damage to the seal plate can result in alarm activations and difficulty with activating the device. - The product analysis found the device's jaw opening and closing mechanism was functioning properly. - The weld integrity of the device was inspected and was found to be within specification. - The blade moved smoothly along the knife track and returned to the home position when the trigger was released. - The knife cut of the device was tested on a silicone test strip with acceptable results. - The heel gap of the device was measured and it was found to be within specification. - The device was activated multiple times on a saline soaked gauze pad but had re-grasp alarms. - The seal cycle was interrupted. - No electrical or wiring issues were found. - Manufacturing does 100% inspection during the assembly and packaging process. The defective sample would have been rejected if found prior to shipping, if this was an assembly defect. The most likely cause was traced to a user issue. The IFU states: Do not attempt to seal or cut over clips or staples as incomplete seals will be formed. The reported condition was confirmed. The analysis noted evidence that the device was not used as intended. The IFU states: Do not attempt to seal or cut over clips or staples as incomplete seals will be formed. Medtronic's assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic's complaint monitoring and tracking system. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Product Id: VLFT10GEN Serial No: UNKNOWN Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all received information. Without a product returned for evaluation a likely cause for the reported condition could not be established. Medtronic's assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic's complaint monitoring and tracking system. The manufacturing records for each device are thoroughly reviewed to ensure the product meets its quality specifications. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical