Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

February 4, 2026

Paul Shaw
101 Manthorpe Road
Grantham NG31 8DG

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| 7209820 | 7209820 TRUCLEAR FOOTSWITCH | SN: YG2572 |
| 7209808 | 7209808 TRUCLEAR CONTROL UNIT | SN: YF7484 |

708614338

Event date: 2025-10-01

Dear Paul Shaw,

Thank you for informing Medtronic of your experience with the above referenced products. The information we received stated during testing, the right foot pedal was unresponsive. The engineer set up the unit with a handpiece and operated forwards/backwards and oscillate lock at different speeds, then kept the forward pedal pressed while moving the foot pedal cable and the handpiece suddenly sped up from 1000 to 2500 and would not stop. The unit had to be turned off to reset it. The issue occurred twice but was very intermittent. There were no error codes displayed. There was no patient involvement. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: 7209820 Serial No: YG2572 Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all received information. Without a product returned for evaluation a likely cause for the reported condition could not be established. Medtronic's assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic's complaint monitoring and tracking system. The manufacturing records for each device are thoroughly reviewed to ensure the product meets its quality specifications. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Product Id: 7209808 Serial No: YF7484 Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all received information. Without a product returned for evaluation a likely cause for the reported condition could not be established. Medtronic's assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic's complaint monitoring and tracking system. The manufacturing records for each device are thoroughly reviewed to ensure the product meets its quality specifications. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical