Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

February 4, 2026

Phil Le Page
PO BOX 958 / NORTH SYDNEY NSW 2059

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| LXMJ37S | MARYLAND XP INLINE SEALER/DIVIDER 37CM | LN: 31460044X |

708532735

MPXR Report # 1336322

Event date: 2025-08-21

Dear Phil Le Page,

Thank you for informing Medtronic of your experience with the above referenced product
      . The information we received stated during an esophagectomy procedure involving the LigaSure Vessel Sealing device (MARYLAND XP INLINE SEALER/DIVIDER 37CM), an energy activation and sealing issue occurred, specifically no seal was achieved, there was no evidence of energy delivery, and end tones were heard. There was also seal cycle error. Bleeding was observed while sealing tissue at the greater curvature of the stomach, which was described as moderate in thickness. The jaws of the device were cleaned quite often during the procedure, and approximately 15-20 successful seals were performed before the issue occurred. Another LigaSure device was used successfully with the same generator. Blood transfusion was not necessary. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: LXMJ37S Lot No: 31460044X Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all received information. Without a product returned for evaluation a likely cause for the reported condition could not be established. Medtronic’s assessment determined that manufacturing action is not required at this time. To ensure product oversight, this complaint report is incorporated into Medtronic’s complaint monitoring and tracking system. The manufacturing records for each device are thoroughly reviewed to ensure the product meets its quality specifications. Should new information become available the file will be re-opened and the investigation summary will be amended as appropriate.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical