Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

November 14, 2025

Ramez Bassari
Epworth Richmond
Bridge Rd Richmond

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| LF1937 | JAW LAP LF1937 LIGASURE MARYLAND 37CM | LN: 43040299X |

707556436

MPXR Report# 1269930

Event date: 2025-02-05

Dear Ramez Bassari,

Thank you for informing Medtronic of your experience with the above referenced product
      . The information we received stated during procedure, the device had inadequate or partial sealing. There was evidence of energy delivery and end tones was heard. There was no bleeding after cutting. The device would not seal from the start, so jaws were cleaned. Another LigaSure device was used successfully with this generator. There was no patient injury. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: LF1937 Lot No: 43040299X JAW LAP LF1937 LIGASURE MARYLAND 37CM was received for evaluation. Examination of the sample is described below. Visual inspection: - The visual inspection found no notable conditions. Evaluation: - The product analysis found the device's jaw opening and closing mechanism was functioning properly. - The weld integrity of the device was inspected and was found to be within specification. - The knife cut of the device was tested on a silicone test strip with acceptable results. - The heel gap of the device was measured and it was found to be within specification. - The device was detected and activated multiple times on a saline soaked gauze pad with satisfactory results. - All seal cycles were completed satisfactorily and end tones were heard indicating completed activation cycles. - The device was activated with the rotation knob in various positions to detect any problematic activation issues. - The device sealing performance was tested on porcine kidney tissue. Multiple seals on various size vessels were made and a full thermal effect was visually verified. The sealing was adequate when the seal cycle reached end tone. - No electrical or wiring issues were found. Based on the evidence available, the reported condition of PARTIAL SEAL was not confirmed. Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during procedure, the device had inadequate or partial sealing. There was evidence of energy delivery and end tones was heard. There was no bleeding after cutting. The device would not seal from the start, so jaws were cleaned. One JAW LAP LF1937 LIGASURE MARYLAND 37CM device was returned for analysis. The device was detected and activated multiple times on a saline soaked gauze pad with satisfactory results. All seal cycles were completed satisfactorily and end tones were heard indicating completed activation cycles. The device sealing performance was tested on porcine kidney tissue. Multiple seals on various size vessels were made and a full thermal effect was visually verified. The sealing was adequate when the seal cycle reached end tone. Based on the evidence available the reported condition(s) could not be confirmed. The most likely cause was not available because no problem was detected with the device. A Device History Record (DHR) or Service History Record (SHR) review is not required since there is no indication of a potential manufacturing or servicing issue. The complaint investigation determined that the device tested within specification/performed as intended. Therefore, no Further Action is needed.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical