Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

November 12, 2025

Amit Katz

Rambam Medical Center
P O B 9602 / 31096 Haifa

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| EEAORVIL21A | EEAORVIL21A EEA ORVIL 21MM AUTOMATICDV | Unknown |
| EEA2135 | EEA2135 EEA 21MM-3.5 SGL USE STAPLER | Unknown |

707589947

MPXR Report# 1271409

RB Reference # {{ref number}}

Event date: 2025-02-11

Dear Amit Katz,

Thank you for informing Medtronic of your experience with the above referenced product{{s}}. The information we received stated during an esophagectomy procedure, there were issues with the OrVil (Oral Anvil) and EEA stapler. There were set-up issues prior to firing, and the oral anvil was difficult to attach to the stapler. The procedure was extended by up to 30 minutes as the devices did not click inside the cavity and were tried outside the patient. The devices were replaced with new ones, which functioned correctly. There was no patient injury.. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: EEAORVIL21A Lot No: Unknown If no Product return “Information provided to Medtronic indicated that the complaint device was not available for evaluation.” Else: {{# of same item}} EEAORVIL21A EEA ORVIL 21MM AUTOMATICDV was received for evaluation. Examination of the sample is described below. 821040768 This report is based on information provided by Returned Product Analysis (RPA) Lab investigation personnel. RPA Lab received one EEAORVIL21A opened by the account without the packaging. Visual inspection: {Anvil} - The EEA2135 instrument handle received under line item 707589947-20 can not be used with an EEAORVIL21A. - The anvil of the instrument was observed to be tilted and separated from the instrument. - Further inspection of the anvil cutting ring showed no traces of knife impression and the ring was received intact. Evaluation: RPA EEAXL instrument used.P6D0308KX - The wing nut and safety functioned properly upon rotating the twist knob. - The green bar was visible within the indicator window when the twist knob was fully closed. - The instrument was applied to the appropriate test media producing acceptable results. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly. - The device also produced all audible, tactile and visual indicators during the functional testing. Based on the evidence available the reported condition of difficult to load was not confirmed. 821040812 This report is based on information provided by Returned Product Analysis (RPA) Lab investigation personnel. RPA Lab received one EEAXL2135 opened by the account without the packaging. Visual inspection: {Instrument} - The visual inspection of the staple guide noted the presence of a full complement of staples. - Further inspection noted that the green bar was not visible in the indicator window and the safety feature was engaged. - The staple guide was observed to be attached to the instrument with no damage to the staple guide. Evaluation: RPN anvil used for testing P0G0155KY - The wing nut and safety functioned properly upon rotating the twist knob. - The green bar was visible within the indicator window when the twist knob was fully closed. - The instrument was applied to the appropriate test media producing acceptable results. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly and completely and a full complement of staples was deployed and properly formed. - The device also produced all audible, tactile and visual indicators during the functional testing. - The instrument body label was removed for evaluation of the eccentric washer assembly. - Evaluation of the Eccentric Washer noted that it was secured. Based on the evidence available the reported condition of orvil anvil difficult to attach was not confirmed. 818975059 Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during an esophagectomy procedure, there were issues with the OrVil (Oral Anvil) and EEA stapler. There were set-up issues prior to firing, and the oral anvil was difficult to attach to the stapler. The procedure was extended by up to 30 minutes as the devices did not click inside the cavity and were tried outside the patient. The devices were replaced with new ones, which functioned correctly. There was no patient injury. One EEAORVIL21A EEA ORVIL 21MM AUTOMATICDV device was returned for analysis. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly. - The device also produced all audible, tactile and visual indicators during the functional testing. Based on the evidence available the reported condition(s) could not be confirmed. The most likely cause was not available because no problem was detected with the device. A Device History Record (DHR) or Service History Record (SHR) review is not required since there is no indication of a potential manufacturing or servicing issue. The complaint investigation determined that the device tested within specification/performed as intended. Therefore, no Further Action is needed. 818982593 Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during an esophagectomy procedure, there were issues with the OrVil (Oral Anvil) and EEA stapler. There were set-up issues prior to firing, and the oral anvil was difficult to attach to the stapler. The procedure was extended by up to 30 minutes as the devices did not click inside the cavity and were tried outside the patient. The devices were replaced with new ones, which functioned correctly. There was no patient injury. One EEA2135 EEA 21MM-3.5 SGL USE STAPLER device was returned for analysis. - The device also produced all audible, tactile and visual indicators during the functional testing. - The instrument body label was removed for evaluation of the eccentric washer assembly. - Evaluation of the Eccentric Washer noted that it was secured. Based on the evidence available the reported condition(s) could not be confirmed. The most likely cause was not available because no problem was detected with the device. A Device History Record (DHR) or Service History Record (SHR) review is not required since there is no indication of a potential manufacturing or servicing issue. The complaint investigation determined that the device tested within specification/performed as intended. Therefore, no Further Action is needed.

Product Id: EEA2135 Serial No: Unknown Information provided to Medtronic indicated that the complaint device was not available for evaluation. 821040768 This report is based on information provided by Returned Product Analysis (RPA) Lab investigation personnel. RPA Lab received one EEAORVIL21A opened by the account without the packaging. Visual inspection: {Anvil} - The EEA2135 instrument handle received under line item 707589947-20 can not be used with an EEAORVIL21A. - The anvil of the instrument was observed to be tilted and separated from the instrument. - Further inspection of the anvil cutting ring showed no traces of knife impression and the ring was received intact. Evaluation: RPA EEAXL instrument used.P6D0308KX - The wing nut and safety functioned properly upon rotating the twist knob. - The green bar was visible within the indicator window when the twist knob was fully closed. - The instrument was applied to the appropriate test media producing acceptable results. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly. - The device also produced all audible, tactile and visual indicators during the functional testing. Based on the evidence available the reported condition of difficult to load was not confirmed. 821040812 This report is based on information provided by Returned Product Analysis (RPA) Lab investigation personnel. RPA Lab received one EEAXL2135 opened by the account without the packaging. Visual inspection: {Instrument} - The visual inspection of the staple guide noted the presence of a full complement of staples. - Further inspection noted that the green bar was not visible in the indicator window and the safety feature was engaged. - The staple guide was observed to be attached to the instrument with no damage to the staple guide. Evaluation: RPN anvil used for testing P0G0155KY - The wing nut and safety functioned properly upon rotating the twist knob. - The green bar was visible within the indicator window when the twist knob was fully closed. - The instrument was applied to the appropriate test media producing acceptable results. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly and completely and a full complement of staples was deployed and properly formed. - The device also produced all audible, tactile and visual indicators during the functional testing. - The instrument body label was removed for evaluation of the eccentric washer assembly. - Evaluation of the Eccentric Washer noted that it was secured. Based on the evidence available the reported condition of orvil anvil difficult to attach was not confirmed. 818975059 Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during an esophagectomy procedure, there were issues with the OrVil (Oral Anvil) and EEA stapler. There were set-up issues prior to firing, and the oral anvil was difficult to attach to the stapler. The procedure was extended by up to 30 minutes as the devices did not click inside the cavity and were tried outside the patient. The devices were replaced with new ones, which functioned correctly. There was no patient injury. One EEAORVIL21A EEA ORVIL 21MM AUTOMATICDV device was returned for analysis. - Pusher-fingers and knife advanced when the handle was squeezed and the knife cut the media cleanly. - The device also produced all audible, tactile and visual indicators during the functional testing. Based on the evidence available the reported condition(s) could not be confirmed. The most likely cause was not available because no problem was detected with the device. A Device History Record (DHR) or Service History Record (SHR) review is not required since there is no indication of a potential manufacturing or servicing issue. The complaint investigation determined that the device tested within specification/performed as intended. Therefore, no Further Action is needed. 818982593 Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during an esophagectomy procedure, there were issues with the OrVil (Oral Anvil) and EEA stapler. There were set-up issues prior to firing, and the oral anvil was difficult to attach to the stapler. The procedure was extended by up to 30 minutes as the devices did not click inside the cavity and were tried outside the patient. The devices were replaced with new ones, which functioned correctly. There was no patient injury. One EEA2135 EEA 21MM-3.5 SGL USE STAPLER device was returned for analysis. - The device also produced all audible, tactile and visual indicators during the functional testing. - The instrument body label was removed for evaluation of the eccentric washer assembly. - Evaluation of the Eccentric Washer noted that it was secured. Based on the evidence available the reported condition(s) could not be confirmed. The most likely cause was not available because no problem was detected with the device. A Device History Record (DHR) or Service History Record (SHR) review is not required since there is no indication of a potential manufacturing or servicing issue. The complaint investigation determined that the device tested within specification/performed as intended. Therefore, no Further Action is needed.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical