Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

November 14, 2025

Carlos Barroso
SERVICLINICA S.A.
LATORRE 98
2290000 CALERA

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| ONB5STF | ONB5STF VERSAONE OPT 5MM STD TROCAR | LN: J4D3663Y |

708143041

MPXR Report# 1304590

Event date: 2025-05-17

Dear Carlos Barroso,

Thank you for informing Medtronic of your experience with the above referenced product
      . The information we received stated during the procedure, at the beginning of the surgery, the valve began to leak air/gas from the seal. The trocar was replaced with another trocar to resolve the issue. There was no patient injury. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: ONB5STF Lot No: J4D3663Y Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during the procedure, at the beginning of the surgery, the valve began to leak air/gas from the seal. The trocar was replaced with another trocar to resolve the issueThe product sample or additional supporting materials from the account were not available for analysis. Based on the evidence available there was not enough information to make any determination. Without the sample a detailed investigation could not be performed, and definitive cause could not be identified. The suspected or most likely cause of the event could not be determined. A review of the Device History Record indicates the product was released meeting all quality release specifications at the time of manufacture. Further action was not required because the event had foreseen risk and is included in a data monitoring plan.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical