Medical Surgical

15 Hampshire Street

Mansfield MA 02048

United States

www.medtronic.com

November 14, 2025

Customer
Anvisa
SIA Trecho 5, Área Especial 57 / Brasilia

In Reference to:

| Product Id | Product Description | Serial No/Lot No |
| --- | --- | --- |
| SIGPSHELL | SIG POWER SIGPSHELL CONTROL SHELL | LN: N3K2219Y |
| SIGPHANDLE | SIG POWER SIGPHANDLE HANDLE | SN: UNKNOWN |

707967075

MPXR Report# 1295384

RB Reference # 2024.08.004129

Event date: 2024-08-21

Dear Customer,

Thank you for informing Medtronic of your experience with the above referenced products. The information we received stated during a surgical procedure, there was a reading failure of the disposable trigger at the time of assembling the powered handle. The rods and battery were replaced due to the persistent failure, but the issue remained unresolved, necessitating the opening of a new trigger. We, as a manufacturer, strive for excellence in constantly improving our products and the service we provide. Customer feedback is very important to our goal of manufacturing and distributing high quality products.

Conclusion:

Product Id: SIGPSHELL Lot No: N3K2219Y Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during a surgical procedure, there was a reading failure of the disposable trigger at the time of assembling the powered handle. The rods and battery were replaced due to the persistent failure, but the issue remained unresolved, necessitating the opening of a new trigger.The product sample or additional supporting materials from the account were not available for analysis. Based on the evidence available there was not enough information to make any determination. Without the sample a detailed investigation could not be performed, and definitive cause could not be identified. The suspected or most likely cause of the event could not be determined. A review of the Device History Record indicates the product was released meeting all quality release specifications at the time of manufacture. Further action was not required because the event had foreseen risk and is included in a data monitoring plan.

Product Id: SIGPHANDLE Serial No: UNKNOWN Information provided to Medtronic indicated that the complaint device was not available for evaluation. Medtronic conducted an investigation based upon all information received. Medtronic was notified of the following reported condition - during a surgical procedure, there was a reading failure of the disposable trigger at the time of assembling the powered handle. The rods and battery were replaced due to the persistent failure, but the issue remained unresolved, necessitating the opening of a new trigger.The product sample or additional supporting materials from the account were not available for analysis. Based on the evidence available there was not enough information to make any determination. Without the sample a detailed investigation could not be performed, and definitive cause could not be identified. The suspected or most likely cause of the event could not be determined. A Lot/Serial Number was not provided or was invalid; therefore, a review of the appropriate Device History Record could not be performed. Further action was not required because the event had foreseen risk and is included in a data monitoring plan.

Thank you for your time and concern in reporting this event. We appreciate the opportunity to evaluate product performance, as this is a valuable source of information for improving reliability and design. If we may be of further assistance, please contact your Medtronic Representative and they will work with our team.

Sincerely,

Tracy Landers

Sr MDR/Vigilance Manager | Customer Quality

Medtronic – Medical Surgical | Surgical